| TYPE A UNCERTAINTY | VALIDATION EMAIL DRAFT |
| --- | --- |

| TO | [Recipient(s)] |
| --- | --- |
| SUBJECT | Validation of Type A uncertainty and expanded reporting |

Hi [Name],

We reviewed the AC-Shunt Type A calculation and its user-facing uncertainty reporting against the GUM and NIST Technical Note 1297. The implemented approach is correct and keeps the quantities used for uncertainty propagation distinct from the value presented as expanded uncertainty.

Calculation approach

- The paired cycle deltas are treated as repeated observations of the measurand.

- Their sample standard deviation uses the unbiased N - 1 denominator.

- The Type A standard uncertainty of the reported mean is u_A = s / sqrt(N). This k = 1 value is retained for uncertainty propagation.

- For concise reporting, the conventional expanded value is U_A = 2u_A. This is an approximate 95% coverage statement under the usual assumptions; it is not labeled as an exact finite-sample confidence interval.

Run example (N = 15)

| Sample standard deviation, s | 3.9797 ppm |
| --- | --- |
| Type A standard uncertainty, u_A | 1.0276 ppm |
| Expanded Type A uncertainty, U_A (k = 2) | 2.0551 ppm |

Qualification: if an exact two-sided 95% confidence interval for a small sample is required, use the Student-t factor. For N = 15 (14 degrees of freedom), t = 2.1448 and the half-width is 2.2039 ppm. That interval answers a different reporting question from the conventional k = 2 expanded uncertainty used here.

Authoritative references

- JCGM 100:2008 (GUM), sections 4.2.2-4.2.3

- NIST TN 1297, Section 3 - Type A Evaluation

- NIST TN 1297, Section 6 - Expanded Uncertainty

- NIST/SEMATECH e-Handbook - t-based confidence interval

Accordingly, the application should display both values in one Type A summary: the standard uncertainty (k = 1) used by the engine and the expanded uncertainty (k = 2) used for concise reporting.

Regards,

[Your Name]